## Title Sheet
Comparative validation of amperometric and optical analyzers of dissolved oxygen: a case study
I. Helm, G. Karina, L. Jalukse, T. Pagano, I. Leito, Environmental Monitoring and Assessment 2018, 190, 313
Irja Helm,1 Gerli Karina,1,2 Lauri Jalukse,1,2 Todd Pagano,3 Ivo Leito1,*
1 University of Tartu, Institute of Chemistry, Ravila 14a, Tartu 50411, Estonia
2 Current address: Estonian Veterinary and Food Laboratory, Kreutzwaldi 30, 51006 Tartu, Estonia
3 Department of Science & Mathematics, Rochester Institute of Technology, Rochester, New York 14623, United States
* Corresponding author: Tel: +372-5 184 176, e-mail: ivo.leito@ut.ee

## Table S1
Table S1 Characteristics of the sensors as provided by manufacturers.
Characteristic | WTW | HACH
Measuring range (linear range) | 0 – 50 mg/L | 0.1 – 20 mg/L
Influence of flow rate | 10% (flow rate) > 3 cm/s) | Not specified
5% (flow rate) > 10 cm/s)
1% (flow rate) > 18 cm/s)
Parameters related to precision and accuracy | ±0,5% (in the range 5-30°C) | ±0.1 mg/L for 0 to 8 mg/L
±0.2 mg/L for greater than 8 mg/L
Accuracy of temperature measurement | ±0.1 °C | ±0.3 °C
Stabilization time | t90 < 10 s | Stabilization time T90% at 10 seconds (when stirred)
t95 < 16 s
t99 < 60 s
Zero signal | < 0,1% | Not specified
Drift (in one month) | 3% | -
Pressure measurement accuracy | - | ±0,8%
Working life without maintenance | min. 6 months with one electrolyte fill | 1 year
Oxygen own consumption at 20 °C | 0.008 μg/h-1 (mg/L)-1 | No own consumption
Source of data | Ref 1 | Ref 2

## Table S2
Table S2. Measurement uncertainty using the ISO GUM method with Kragten approach
Data from River Emajõgi
HACH
Experimental data for calculating different uncertainty contributions:
u(Ccal_rep) | 0.0039025 | mg/L | Pooled standard deviation of the readings at stirring rate 120 rpm
u(Ccal_0_rep) | 0.004913 | mg/L | The largest pooled standard deviation of the readings
u(Cmeas_Rep) | 0.068644 | mg/L | Reproducibility: pooled standard deviation of differences between reading and reference value
u(Ccal) | 0.031 | mg/L | From Ref 3
b1stir | -0.00094245 | %/% | The slope of stirring graph (% of the reading vs % of the stirring rate from calibration stirring rate)
b1Δt | 0.024369 | %/°C | The slope of temperature compensation graph (temperature °C vs relative difference from reference temperature in %)
b1drift | -0.0078995 | mg/L | The slope of the drift graph (number of days passed from calibration vs reading difference from the reference value mg/L)
s(b1drift) | 0.00084931 | mg/L | Standard deviation of the slope of the drift graph
b1linearity | 0.027665 | - | The slope of the linearity graph starting at 11 mg/L (electrochemical reading mg/L vs readings difference mg/L)
Input quantities
| Value | Unit | Comment
Calibration
Cread_cal_0 | 0 | mg/L | Sensor reading right after calibration/checking zero value
Cread_cal | 8.95 | mg/L | Sensor reading right after calibration | [7 empty] | Constants for salinity calculation (Ref 6):
Ccal_0 | 0 | mg/L | Calibration environment DO concentration at oxygen-free conditions, here always 0 | [7 empty] | a0 | 0.008
Ccal | 8.8907 | mg/L | Calibration envirnment DO concentration at saturation, calculated by the Ref 4 | [7 empty] | a1 | -0.1692 |  | C(KCl at 25°C) *
tcal_ref | 19.929 | °C | Temperature of calibration solution measured with reference thermometer | [7 empty] | a2 | 25.3851 |  | 53.9038 | mS/cm
Pcal_ref | 98991 | Pa | Air pressure at the time of calibration measured with reference barometer | [7 empty] | a3 | 14.0941 |  | *Notes:
tcal_sensor | 20.1 | °C | Temperature of calibration solution measured with DO-meter | [7 empty] | a4 | -7.0261 |  | Ref 5: | conductivity of SSW at SP=35, t_68=15, p=0; 42.9140 mS/cm
Pcal_sensor | 100100 | Pa | Air pressure at the time of calibration measured with DO-meter | [7 empty] | a5 | 2.7081 |  | Ref 7: | factor for calculating the conductivity from 15 °C to the 25 °C is found as 1.25608904610271
[11 empty] | b0 | 0.0005
Measurement | [10 empty] | b1 | -0.0056
Cread_meas | 9.85 | mg/L | Reading of the measured sample | [7 empty] | b2 | -0.0066
k | 0.34 | mS/cm | conductivity of the sample, used for calculating the activity coefficient | [7 empty] | b3 | -0.0375
Salinity | 0.15885 | PSU | Salinity, calculated from measured conductivity (recalculated to 25 C), using Refs 5, 6 and 7 | [7 empty] | b4 | 0.0636
tsample | 9.6 | °C | Sample temperature, measured with DO-meter | [7 empty] | b5 | -0.0144
Δvstir | 50 | % | estimated flow rate compared to the flow rate of calibration | [7 empty] | X | 400 | *Rt
Δdays | 3 |  | The number of days passed from the last calibration | [7 empty] | Y | 100 | *Rt
Model components |  |  |  |  | Uncertainty components
Denotion | Value | Unit | uc | Relative | u1 |  |  | u2
Cread_cal_0 | 0 | mg/L | 0.0056984 |  | 0.0028868 | u(Ccal_res) | ±half of the last digit | 0.004913 | u(Ccal_rep) | repeatabilty of the reading
Cread_cal | 8.95 | mg/L | 0.0048542 | 0.05424% | 0.0028868 | u(Ccal_0_res) | ±half of the last digit | 0.0039025 | u(Ccal_0_rep) | repeatabilty of the reading
Ccal_0 | 0 | mg/L | 0.0057735 |  | 0.0057735 | u(Ccal_0) | estimation
Ccal | 8.8907 | mg/L | 0.031 | 0.3487% | 0.031 | u(Ccal) | Estimated as in Ref 3
Cread_meas | 9.85 | mg/L | 0.068705 | 0.6975% | 0.0028868 | u(Cmeas_res) | ±half of the last digit | 0.068644 | u(Cmeas_Rep) | Reproducibility of the reading
γ | 1.001 | - | 0.0002944 | 0.02941% | 0.0002944 | ± half of the activity coefficient difference from 1
ΔCstir | 0 | mg/L | 0.0026798 |  | 0.0026798 | ±Δvstir *b1stir*Cread_meas/100%
ΔCΔt | 0 | mg/L | 0.014551 |  | 0.014551 | ±Δt*b1Δt/100%*Cread_meas
ΔCdrift | 0 | mg/L | 0.015153 |  | 0.015153 | ±Δdays*(b1drift+s(b1drift))
ΔClinearity | 0 | mg/L | 0 |  | 0 | ±ΔC from 11 mg/L*b1linearity (if Cread_meas>11)
Denotion | Value | Unit | uc |  |  | Cread_cal_0 | Cread_cal | Ccal_0 | Ccal | Cread_meas | γ | ΔCstir | ΔCΔt | ΔCdrift | ΔClinearity
Cread_cal_0 | 0 | mg/L | 0.0056984 |  |  | 0.0056984 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cread_cal | 8.95 | mg/L | 0.0048542 |  |  | 8.95 | 8.9549 | 8.95 | 8.95 | 8.95 | 8.95 | 8.95 | 8.95 | 8.95 | 8.95
Ccal_0 | 0 | mg/L | 0.0057735 |  |  | 0 | 0 | 0.0057735 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ccal | 8.8907 | mg/L | 0.031 |  |  | 8.8907 | 8.8907 | 8.8907 | 8.9217 | 8.8907 | 8.8907 | 8.8907 | 8.8907 | 8.8907 | 8.8907
Cread_meas | 9.85 | mg/L | 0.068705 |  |  | 9.85 | 9.85 | 9.85 | 9.85 | 9.9187 | 9.85 | 9.85 | 9.85 | 9.85 | 9.85
γ | 1.001 | - | 0.0002944 |  |  | 1.001 | 1.001 | 1.001 | 1.001 | 1.001 | 1.0013 | 1.001 | 1.001 | 1.001 | 1.001
ΔCstir | 0 | mg/L | 0.0026798 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.0026798 | 0 | 0 | 0
ΔCΔt | 0 | mg/L | 0.014551 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.014551 | 0 | 0
ΔCdrift | 0 | mg/L | 0.015153 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.015153 | 0
ΔClinearity | 0 | mg/L | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cresult | 9.7748 | mg/L |  |  |  | 9.7753 | 9.7695 | 9.7742 | 9.8088 | 9.8429 | 9.7719 | 9.7774 | 9.7892 | 9.7898 | 9.7748
uCresult | 0.079302 | mg/L |  |  |  | 0.00056323 | -0.0052986 | -0.0005741 | 0.034083 | 0.06818 | -0.0028739 | 0.0026593 | 0.01444 | 0.015038 | 0
UCresult | 0.1586 | mg/L, k=2 |  |  |  | 3.1723e-07 | 2.8076e-05 | 3.2959e-07 | 0.0011616 | 0.0046485 | 8.2594e-06 | 7.072e-06 | 0.00020851 | 0.00022613 | 0
UCresult rel | 1.6226 | %, k=2 | [12 empty] | 0.0062888
| En (vs WTW) = | -0.76396 |  |  |  | 0.005044% | 0.4464% | 0.005241% | 18.47% | 73.92% | 0.1313% | 0.1125% | 3.316% | 3.596% | 0%
100%

## Table S3
Table S3. Measurement uncertainty using the ISO GUM method with Kragten approach
Data from River Emajõgi
WTW
Experimental data for calculating different uncertainty contributions:
u(Ccal_rep) | 0.019833 | mg/L | Pooled standard deviation of the readings at stirring rate 120 rpm
u(Ccal_0_rep) | 0.019833 | mg/L | The largest pooled standard deviation of the readings
u(Cmeas_Rep) | 0.051188 | mg/L | Reproducibility: pooled standard deviation of differences between reading and reference value
u(Ccal) | 0.031 | mg/L | From Ref 3
b1stir | 0.014614 | %/% | The slope of stirring graph (% of the reading vs % of the stirring rate from calibration stirring rate)
b1Δt | -0.077609 | %/°C | The slope of temperature compensation graph (temperature °C vs relative difference from reference temperature in %)
b1drift | 0.0037617 | mg/L | The largest slope of the drift graph (number of days passed from calibration vs reading difference from the reference value mg/L)
s(b1drift) | 0.00089586 | mg/L | Standard deviation of the slope of the drift graph
Input quantities
| Value | Unit | Comment
Calibration
Cread_cal_0 | 0 | mg/L | Sensor reading right after calibration/checking zero value | [7 empty] | Constants for salinity calculation (Ref 6):
Cread_cal | 8.89 | mg/L | Sensor reading right after calibration | [7 empty] | a0 | 0.008
Ccal_0 | 0 | mg/L | Calibration environment DO concentration at oxygen-free conditions, here always 0 | [7 empty] | a1 | -0.1692
Ccal | 8.8907 | mg/L | Calibration envirnment DO concentration at saturation, calculated by the Ref 4 | [7 empty] | a2 | 25.3851 |  | C(KCl at 25°C) *
tcal_ref | 19.929 | °C | Temperature of calibration solution measured with reference thermometer | [7 empty] | a3 | 14.0941 |  | 53.9038 | mS/cm
Pcal_ref | 98991 | Pa | Air pressure at the time of calibration measured with reference barometer | [7 empty] | a4 | -7.0261 |  | *Notes:
tcal_sensor | 19.9 | °C | Temperature of calibration solution measured with DO-meter | [7 empty] | a5 | 2.7081 |  | Ref 5: | conductivity of SSW at SP=35, t_68=15, p=0; 42.9140 mS/cm
Pcal_sensor | 99200 | Pa | Air pressure at the time of calibration measured with DO-meter | [7 empty] | b0 | 0.0005 |  | Ref 7: | factor for calculating the conductivity from 15 °C to the 25 °C is found as 1.25608904610271
[11 empty] | b1 | -0.0056
Measurement | [10 empty] | b2 | -0.0066
Cread_meas | 9.97 | mg/L | Reading of the measured sample | [7 empty] | b3 | -0.0375
k | 0.34 | mS/cm | conductivity of the sample, used for calculating the activity coefficient | [7 empty] | b4 | 0.0636
Salinity | 0.15885 | PSU | Salinity, calculated from measured conductivity (recalculated to 25 C), using Refs 5, 6 and 7 | [7 empty] | b5 | -0.0144
tsample | 9.1 | °C | Sample temperature, measured with DO-meter | [7 empty] | X | 400 | *Rt
Δvstir | 50 | % | estimated flow rate compared to the flow rate of calibration | [7 empty] | Y | 100 | *Rt
Δdays | 3 |  | The number of days passed from the last calibration
Model components |  |  |  |  | Uncertainty components
Denotion | Value | Unit | uc | Relative | u1 |  |  | u2
Cread_cal_0 | 0 | mg/L | 0.020042 |  | 0.0028868 | u(Ccal_res) | ±half of the last digit | 0.019833 | u(Ccal_rep) | repeatabilty of the reading
Cread_cal | 8.89 | mg/L | 0.020042 | 0.2254% | 0.0028868 | u(Ccal_0_res) | ±half of the last digit | 0.019833 | u(Ccal_0_rep) | repeatabilty of the reading
Ccal_0 | 0 | mg/L | 0.0057735 |  | 0.0057735 | u(Ccal_0) | estimation
Ccal | 8.8907 | mg/L | 0.031 | 0.3487% | 0.031 | u(Ccal) | Estimated as in Ref 3
Cread_meas | 9.97 | mg/L | 0.051269 | 0.5142% | 0.0028868 | u(Cmeas_res) | ±half of the last digit | 0.051188 | u(Cmeas_Rep) | reproducibility of the reading
γ | 1.001 | - | 0.00029567 | 0.02954% | 0.00029567 | ± half of the activity coefficient difference from 1
ΔCstir | 0 | mg/L | 0.042061 |  | 0.042061 | ±Δvstir *b1stir*Cread_meas/100%
ΔCΔt | 0 | mg/L | 0.048247 |  | 0.048247 | ±Δt*b1Δt/100%*Cread_meas
ΔCdrift | 0 | mg/L | 0.0080672 |  | 0.0080672 | ±Δdays*(b1drift+s(b1drift))
Denotion | Value | Unit | uc |  |  | Cread_cal_0 | Cread_cal | Ccal_0 | Ccal | Cread_meas | γ | ΔCstir | ΔCΔt | ΔCdrift
Cread_cal_0 | 0 | mg/L | 0.020042 |  |  | 0.020042 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cread_cal | 8.89 | mg/L | 0.020042 |  |  | 8.89 | 8.91 | 8.89 | 8.89 | 8.89 | 8.89 | 8.89 | 8.89 | 8.89
Ccal_0 | 0 | mg/L | 0.0057735 |  |  | 0 | 0 | 0.0057735 | 0 | 0 | 0 | 0 | 0 | 0
Ccal | 8.8907 | mg/L | 0.031 |  |  | 8.8907 | 8.8907 | 8.8907 | 8.9217 | 8.8907 | 8.8907 | 8.8907 | 8.8907 | 8.8907
Cread_meas | 9.97 | mg/L | 0.051269 |  |  | 9.97 | 9.97 | 9.97 | 9.97 | 10.0213 | 9.97 | 9.97 | 9.97 | 9.97
γ | 1.001 | - | 0.00029567 |  |  | 1.001 | 1.001 | 1.001 | 1.001 | 1.001 | 1.0013 | 1.001 | 1.001 | 1.001
ΔCstir | 0 | mg/L | 0.042061 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.042061 | 0 | 0
ΔCΔt | 0 | mg/L | 0.048247 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.048247 | 0
ΔCdrift | 0 | mg/L | 0.0080672 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0080672
Cresult | 9.9606 | mg/L |  |  |  | 9.963 | 9.9382 | 9.9599 | 9.9953 | 10.0118 | 9.9576 | 10.0026 | 10.0088 | 9.9686
uCresult | 0.0922 | mg/L |  |  |  | 0.0024174 | -0.022405 | -0.00069477 | 0.03473 | 0.051221 | -0.0029411 | 0.042021 | 0.048202 | 0.0080596
UCresult | 0.1844 | mg/L, k=2 |  |  |  | 5.8438e-06 | 0.00050197 | 4.827e-07 | 0.0012062 | 0.0026236 | 8.6502e-06 | 0.0017658 | 0.0023234 | 6.4957e-05
UCresult rel | 1.8513 | %, k=2 | [11 empty] | 0.0085009
[6 empty] | 0.06874% | 5.905% | 0.005678% | 14.19% | 30.86% | 0.1018% | 20.77% | 27.33% | 0.7641%
100%

## Table S4
Table S4. Measurement uncertainty using the ISO GUM method with Kragten approach
Data from Lake Ilmatsalu 1
HACH
Experimental data for calculating different uncertainty contributions:
u(Ccal_rep) | 0.0039025 | mg/L | Pooled standard deviation of the readings at stirring rate 120 rpm
u(Ccal_0_rep) | 0.004913 | mg/L | The largest pooled standard deviation of the readings
u(Cmeas_Rep) | 0.068644 | mg/L | Reproducibility: pooled standard deviation of differences between reading and reference value
u(Ccal) | 0.031 | mg/L | From Ref 3
b1stir | -0.00094245 | %/% | The slope of stirring graph (% of the reading vs % of the stirring rate from calibration stirring rate)
b1Δt | 0.024369 | %/°C | The slope of temperature compensation graph (temperature °C vs relative difference from reference temperature in %)
b1drift | -0.0078995 | mg/L | The slope of the drift graph (number of days passed from calibration vs reading difference from the reference value mg/L)
s(b1drift) | 0.00084931 | mg/L | Standard deviation of the slope of the drift graph
b1linearity | 0.027665 | - | The slope of the linearity graph starting at 11 mg/L (electrochemical reading mg/L vs readings difference mg/L)
Input quantities
| Value | Unit | Comment
Calibration
Cread_cal_0 | 0 | mg/L | Sensor reading right after calibration/checking zero value
Cread_cal | 8.95 | mg/L | Sensor reading right after calibration | [7 empty] | Constants for salinity calculation (Ref 6):
Ccal_0 | 0 | mg/L | Calibration environment DO concentration at oxygen-free conditions, here always 0 | [7 empty] | a0 | 0.008
Ccal | 8.8907 | mg/L | Calibration envirnment DO concentration at saturation, calculated by the Ref 4 | [7 empty] | a1 | -0.1692 |  | C(KCl at 25°C) *
tcal_ref | 19.929 | °C | Temperature of calibration solution measured with reference thermometer | [7 empty] | a2 | 25.3851 |  | 53.9038 | mS/cm
Pcal_ref | 98991 | Pa | Air pressure at the time of calibration measured with reference barometer | [7 empty] | a3 | 14.0941 |  | *Notes:
tcal_sensor | 20.1 | °C | Temperature of calibration solution measured with DO-meter | [7 empty] | a4 | -7.0261 |  | Ref 5: | conductivity of SSW at SP=35, t_68=15, p=0; 42.9140 mS/cm
Pcal_sensor | 100100 | Pa | Air pressure at the time of calibration measured with DO-meter | [7 empty] | a5 | 2.7081 |  | Ref 7: | factor for calculating the conductivity from 15 °C to the 25 °C is found as 1.25608904610271
[11 empty] | b0 | 0.0005
Measurement | [10 empty] | b1 | -0.0056
Cread_meas | 10.94 | mg/L | Reading of the measured sample | [7 empty] | b2 | -0.0066
k | 0.578 | mS/cm | conductivity of the sample, used for calculating the activity coefficient | [7 empty] | b3 | -0.0375
Salinity | 0.27428 | PSU | Salinity, calculated from measured conductivity (recalculated to 25 C), using Refs 5, 6 and 7 | [7 empty] | b4 | 0.0636
tsample | 16 | °C | Sample temperature, measured with DO-meter | [7 empty] | b5 | -0.0144
Δvstir | 50 | % | estimated flow rate compared to the flow rate of calibration | [7 empty] | X | 400 | *Rt
Δdays | 24 |  | The number of days passed from the last calibration | [7 empty] | Y | 100 | *Rt
Model components |  |  |  |  | Uncertainty components
Denotion | Value | Unit | uc | Relative | u1 |  |  | u2
Cread_cal_0 | 0 | mg/L | 0.0056984 |  | 0.0028868 | u(Ccal_res) | ±half of the last digit | 0.004913 | u(Ccal_rep) | repeatabilty of the reading
Cread_cal | 8.95 | mg/L | 0.0048542 | 0.05424% | 0.0028868 | u(Ccal_0_res) | ±half of the last digit | 0.0039025 | u(Ccal_0_rep) | repeatabilty of the reading
Ccal_0 | 0 | mg/L | 0.0057735 |  | 0.0057735 | u(Ccal_0) | estimation
Ccal | 8.8907 | mg/L | 0.031 | 0.3487% | 0.031 | u(Ccal) | Estimated as in Ref 3
Cread_meas | 10.94 | mg/L | 0.068705 | 0.628% | 0.0028868 | u(Cmeas_res) | ±half of the last digit | 0.068644 | u(Cmeas_Rep) | Reproducibility of the reading
γ | 1.0017 | - | 0.00048231 | 0.04815% | 0.00048231 | ± half of the activity coefficient difference from 1
ΔCstir | 0 | mg/L | 0.0029763 |  | 0.0029763 | ±Δvstir *b1stir*Cread_meas/100%
ΔCΔt | 0 | mg/L | 0.0063106 |  | 0.0063106 | ±Δt*b1Δt/100%*Cread_meas
ΔCdrift | 0 | mg/L | 0.12123 |  | 0.12123 | ±Δdays*(b1drift+s(b1drift))
ΔClinearity | 0 | mg/L | 0 |  | 0 | ±ΔC from 11 mg/L*b1linearity (if Cread_meas>11)
Denotion | Value | Unit | uc |  |  | Cread_cal_0 | Cread_cal | Ccal_0 | Ccal | Cread_meas | γ | ΔCstir | ΔCΔt | ΔCdrift | ΔClinearity
Cread_cal_0 | 0 | mg/L | 0.0056984 |  |  | 0.0056984 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cread_cal | 8.95 | mg/L | 0.0048542 |  |  | 8.95 | 8.9549 | 8.95 | 8.95 | 8.95 | 8.95 | 8.95 | 8.95 | 8.95 | 8.95
Ccal_0 | 0 | mg/L | 0.0057735 |  |  | 0 | 0 | 0.0057735 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ccal | 8.8907 | mg/L | 0.031 |  |  | 8.8907 | 8.8907 | 8.8907 | 8.9217 | 8.8907 | 8.8907 | 8.8907 | 8.8907 | 8.8907 | 8.8907
Cread_meas | 10.94 | mg/L | 0.068705 |  |  | 10.94 | 10.94 | 10.94 | 10.94 | 11.0087 | 10.94 | 10.94 | 10.94 | 10.94 | 10.94
γ | 1.0017 | - | 0.00048231 |  |  | 1.0017 | 1.0017 | 1.0017 | 1.0017 | 1.0017 | 1.0022 | 1.0017 | 1.0017 | 1.0017 | 1.0017
ΔCstir | 0 | mg/L | 0.0029763 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.0029763 | 0 | 0 | 0
ΔCΔt | 0 | mg/L | 0.0063106 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0063106 | 0 | 0
ΔCdrift | 0 | mg/L | 0.12123 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.12123 | 0
ΔClinearity | 0 | mg/L | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cresult | 10.8494 | mg/L |  |  |  | 10.8506 | 10.8435 | 10.8481 | 10.8872 | 10.9175 | 10.8441 | 10.8523 | 10.8556 | 10.9696 | 10.8494
uCresult | 0.14367 | mg/L |  |  |  | 0.0012479 | -0.0058812 | -0.0012719 | 0.03783 | 0.068135 | -0.0052215 | 0.0029517 | 0.0062584 | 0.12022 | 0
UCresult | 0.28733 | mg/L, k=2 |  |  |  | 1.5572e-06 | 3.4588e-05 | 1.6178e-06 | 0.0014311 | 0.0046424 | 2.7264e-05 | 8.7125e-06 | 3.9167e-05 | 0.014454 | 0
UCresult rel | 2.6484 | %, k=2 | [12 empty] | 0.02064
| En (vs WTW) = | -2.1735 |  |  |  | 0.007544% | 0.1676% | 0.007838% | 6.934% | 22.49% | 0.1321% | 0.04221% | 0.1898% | 70.03% | 0%
100%

## Table S5
Table S5. Measurement uncertainty using the ISO GUM method with Kragten approach
Data from Lake Ilmatsalu 1
WTW
Experimental data for calculating different uncertainty contributions:
u(Ccal_rep) | 0.019833 | mg/L | Pooled standard deviation of the readings at stirring rate 120 rpm
u(Ccal_0_rep) | 0.019833 | mg/L | The largest pooled standard deviation of the readings
u(Cmeas_Rep) | 0.051188 | mg/L | Reproducibility: pooled standard deviation of differences between reading and reference value
u(Ccal) | 0.031 | mg/L | From Ref 3
b1stir | 0.014614 | %/% | The slope of stirring graph (% of the reading vs % of the stirring rate from calibration stirring rate)
b1Δt | -0.077609 | %/°C | The slope of temperature compensation graph (temperature °C vs relative difference from reference temperature in %)
b1drift | 0.0037617 | mg/L | The largest slope of the drift graph (number of days passed from calibration vs reading difference from the reference value mg/L)
s(b1drift) | 0.00089586 | mg/L | Standard deviation of the slope of the drift graph
Input quantities
| Value | Unit | Comment
Calibration
Cread_cal_0 | 0 | mg/L | Sensor reading right after calibration/checking zero value | [7 empty] | Constants for salinity calculation (Ref 6):
Cread_cal | 8.89 | mg/L | Sensor reading right after calibration | [7 empty] | a0 | 0.008
Ccal_0 | 0 | mg/L | Calibration environment DO concentration at oxygen-free conditions, here always 0 | [7 empty] | a1 | -0.1692
Ccal | 8.8907 | mg/L | Calibration envirnment DO concentration at saturation, calculated by the Ref 4 | [7 empty] | a2 | 25.3851 |  | C(KCl at 25°C) *
tcal_ref | 19.929 | °C | Temperature of calibration solution measured with reference thermometer | [7 empty] | a3 | 14.0941 |  | 53.9038 | mS/cm
Pcal_ref | 98991 | Pa | Air pressure at the time of calibration measured with reference barometer | [7 empty] | a4 | -7.0261 |  | *Notes:
tcal_sensor | 19.9 | °C | Temperature of calibration solution measured with DO-meter | [7 empty] | a5 | 2.7081 |  | Ref 5: | conductivity of SSW at SP=35, t_68=15, p=0; 42.9140 mS/cm
Pcal_sensor | 99200 | Pa | Air pressure at the time of calibration measured with DO-meter | [7 empty] | b0 | 0.0005 |  | Ref 7: | factor for calculating the conductivity from 15 °C to the 25 °C is found as 1.25608904610271
[11 empty] | b1 | -0.0056
Measurement | [10 empty] | b2 | -0.0066
Cread_meas | 11.655 | mg/L | Reading of the measured sample | [7 empty] | b3 | -0.0375
k | 0.578 | mS/cm | conductivity of the sample, used for calculating the activity coefficient | [7 empty] | b4 | 0.0636
Salinity | 0.27428 | PSU | Salinity, calculated from measured conductivity (recalculated to 25 C), using Refs 5, 6 and 7 | [7 empty] | b5 | -0.0144
tsample | 15.4 | °C | Sample temperature, measured with DO-meter | [7 empty] | X | 400 | *Rt
Δvstir | 50 | % | estimated flow rate compared to the flow rate of calibration | [7 empty] | Y | 100 | *Rt
Δdays | 24 |  | The number of days passed from the last calibration
Model components |  |  |  |  | Uncertainty components
Denotion | Value | Unit | uc | Relative | u1 |  |  | u2
Cread_cal_0 | 0 | mg/L | 0.020042 |  | 0.0028868 | u(Ccal_res) | ±half of the last digit | 0.019833 | u(Ccal_rep) | repeatabilty of the reading
Cread_cal | 8.89 | mg/L | 0.020042 | 0.2254% | 0.0028868 | u(Ccal_0_res) | ±half of the last digit | 0.019833 | u(Ccal_0_rep) | repeatabilty of the reading
Ccal_0 | 0 | mg/L | 0.0057735 |  | 0.0057735 | u(Ccal_0) | estimation
Ccal | 8.8907 | mg/L | 0.031 | 0.3487% | 0.031 | u(Ccal) | Estimated as in Ref 3
Cread_meas | 11.655 | mg/L | 0.051269 | 0.4399% | 0.0028868 | u(Cmeas_res) | ±half of the last digit | 0.051188 | u(Cmeas_Rep) | reproducibility of the reading
γ | 1.0017 | - | 0.00048463 | 0.04838% | 0.00048463 | ± half of the activity coefficient difference from 1
ΔCstir | 0 | mg/L | 0.04917 |  | 0.04917 | ±Δvstir *b1stir*Cread_meas/100%
ΔCΔt | 0 | mg/L | 0.023501 |  | 0.023501 | ±Δt*b1Δt/100%*Cread_meas
ΔCdrift | 0 | mg/L | 0.064538 |  | 0.064538 | ±Δdays*(b1drift+s(b1drift))
Denotion | Value | Unit | uc |  |  | Cread_cal_0 | Cread_cal | Ccal_0 | Ccal | Cread_meas | γ | ΔCstir | ΔCΔt | ΔCdrift
Cread_cal_0 | 0 | mg/L | 0.020042 |  |  | 0.020042 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cread_cal | 8.89 | mg/L | 0.020042 |  |  | 8.89 | 8.91 | 8.89 | 8.89 | 8.89 | 8.89 | 8.89 | 8.89 | 8.89
Ccal_0 | 0 | mg/L | 0.0057735 |  |  | 0 | 0 | 0.0057735 | 0 | 0 | 0 | 0 | 0 | 0
Ccal | 8.8907 | mg/L | 0.031 |  |  | 8.8907 | 8.8907 | 8.8907 | 8.9217 | 8.8907 | 8.8907 | 8.8907 | 8.8907 | 8.8907
Cread_meas | 11.655 | mg/L | 0.051269 |  |  | 11.655 | 11.655 | 11.655 | 11.655 | 11.7063 | 11.655 | 11.655 | 11.655 | 11.655
γ | 1.0017 | - | 0.00048463 |  |  | 1.0017 | 1.0017 | 1.0017 | 1.0017 | 1.0017 | 1.0022 | 1.0017 | 1.0017 | 1.0017
ΔCstir | 0 | mg/L | 0.04917 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.04917 | 0 | 0
ΔCΔt | 0 | mg/L | 0.023501 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.023501 | 0
ΔCdrift | 0 | mg/L | 0.064538 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.064538
Cresult | 11.6364 | mg/L |  |  |  | 11.6426 | 11.6102 | 11.6346 | 11.6769 | 11.6875 | 11.6307 | 11.6855 | 11.6598 | 11.7008
uCresult | 0.11017 | mg/L |  |  |  | 0.0062038 | -0.026174 | -0.001783 | 0.040574 | 0.051187 | -0.0056271 | 0.049091 | 0.023463 | 0.064434
UCresult | 0.22034 | mg/L, k=2 |  |  |  | 3.8488e-05 | 0.00068508 | 3.1791e-06 | 0.0016462 | 0.0026201 | 3.1665e-05 | 0.0024099 | 0.00055051 | 0.0041518
UCresult rel | 1.8935 | %, k=2 | [11 empty] | 0.012137
[6 empty] | 0.3171% | 5.645% | 0.02619% | 13.56% | 21.59% | 0.2609% | 19.86% | 4.536% | 34.21%
100%

## Table S6
Table S6. Measurement uncertainty using the ISO GUM method with Kragten approach
Data from Lake Ilmatsalu 2
HACH
Experimental data for calculating different uncertainty contributions:
u(Ccal_rep) | 0.0039025 | mg/L | Pooled standard deviation of the readings at stirring rate 120 rpm
u(Ccal_0_rep) | 0.004913 | mg/L | The largest pooled standard deviation of the readings
u(Cmeas_Rep) | 0.068644 | mg/L | Reproducibility: pooled standard deviation of differences between reading and reference value
u(Ccal) | 0.031 | mg/L | From Ref 3
b1stir | -0.00094245 | %/% | The slope of stirring graph (% of the reading vs % of the stirring rate from calibration stirring rate)
b1Δt | 0.024369 | %/°C | The slope of temperature compensation graph (temperature °C vs relative difference from reference temperature in %)
b1drift | -0.0078995 | mg/L | The slope of the drift graph (number of days passed from calibration vs reading difference from the reference value mg/L)
s(b1drift) | 0.00084931 | mg/L | Standard deviation of the slope of the drift graph
b1linearity | 0.027665 | - | The slope of the linearity graph starting at 11 mg/L (electrochemical reading mg/L vs readings difference mg/L)
Input quantities
| Value | Unit | Comment
Calibration
Cread_cal_0 | 0 | mg/L | Sensor reading right after calibration/checking zero value
Cread_cal | 9.03 | mg/L | Sensor reading right after calibration | [7 empty] | Constants for salinity calculation (Ref 6):
Ccal_0 | 0 | mg/L | Calibration environment DO concentration at oxygen-free conditions, here always 0 | [7 empty] | a0 | 0.008
Ccal | 8.9792 | mg/L | Calibration envirnment DO concentration at saturation, calculated by the Ref 4 | [7 empty] | a1 | -0.1692 |  | C(KCl at 25°C) *
tcal_ref | 19.909 | °C | Temperature of calibration solution measured with reference thermometer | [7 empty] | a2 | 25.3851 |  | 53.9038 | mS/cm
Pcal_ref | 99915 | Pa | Air pressure at the time of calibration measured with reference barometer | [7 empty] | a3 | 14.0941 |  | *Notes:
tcal_sensor | 20.1 | °C | Temperature of calibration solution measured with DO-meter | [7 empty] | a4 | -7.0261 |  | Ref 5: | conductivity of SSW at SP=35, t_68=15, p=0; 42.9140 mS/cm
Pcal_sensor | 100100 | Pa | Air pressure at the time of calibration measured with DO-meter | [7 empty] | a5 | 2.7081 |  | Ref 7: | factor for calculating the conductivity from 15 °C to the 25 °C is found as 1.25608904610271
[11 empty] | b0 | 0.0005
Measurement | [10 empty] | b1 | -0.0056
Cread_meas | 11.95 | mg/L | Reading of the measured sample | [7 empty] | b2 | -0.0066
k | 0.558 | mS/cm | conductivity of the sample, used for calculating the activity coefficient | [7 empty] | b3 | -0.0375
Salinity | 0.26449 | PSU | Salinity, calculated from measured conductivity (recalculated to 25 C), using Refs 5, 6 and 7 | [7 empty] | b4 | 0.0636
tsample | 19.2 | °C | Sample temperature, measured with DO-meter | [7 empty] | b5 | -0.0144
Δvstir | 50 | % | estimated flow rate compared to the flow rate of calibration | [7 empty] | X | 400 | *Rt
Δdays | 0 |  | The number of days passed from the last calibration | [7 empty] | Y | 100 | *Rt
Model components |  |  |  |  | Uncertainty components
Denotion | Value | Unit | uc | Relative | u1 |  |  | u2
Cread_cal_0 | 0 | mg/L | 0.0056984 |  | 0.0028868 | u(Ccal_res) | ±half of the last digit | 0.004913 | u(Ccal_rep) | repeatabilty of the reading
Cread_cal | 9.03 | mg/L | 0.0048542 | 0.05376% | 0.0028868 | u(Ccal_0_res) | ±half of the last digit | 0.0039025 | u(Ccal_0_rep) | repeatabilty of the reading
Ccal_0 | 0 | mg/L | 0.0057735 |  | 0.0057735 | u(Ccal_0) | estimation
Ccal | 8.9792 | mg/L | 0.031 | 0.3452% | 0.031 | u(Ccal) | Estimated as in Ref 3
Cread_meas | 11.95 | mg/L | 0.068705 | 0.5749% | 0.0028868 | u(Cmeas_res) | ±half of the last digit | 0.068644 | u(Cmeas_Rep) | Reproducibility of the reading
γ | 1.0016 | - | 0.00045357 | 0.04529% | 0.00045357 | ± half of the activity coefficient difference from 1
ΔCstir | 0 | mg/L | 0.0032511 |  | 0.0032511 | ±Δvstir *b1stir*Cread_meas/100%
ΔCΔt | 0 | mg/L | 0.0015132 |  | 0.0015132 | ±Δt*b1Δt/100%*Cread_meas
ΔCdrift | 0 | mg/L | 0 |  | 0 | ±Δdays*(b1drift+s(b1drift))
ΔClinearity | 0 | mg/L | 0.015174 |  | 0.015174 | ±ΔC from 11 mg/L*b1linearity (if Cread_meas>11)
Denotion | Value | Unit | uc |  |  | Cread_cal_0 | Cread_cal | Ccal_0 | Ccal | Cread_meas | γ | ΔCstir | ΔCΔt | ΔCdrift | ΔClinearity
Cread_cal_0 | 0 | mg/L | 0.0056984 |  |  | 0.0056984 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cread_cal | 9.03 | mg/L | 0.0048542 |  |  | 9.03 | 9.0349 | 9.03 | 9.03 | 9.03 | 9.03 | 9.03 | 9.03 | 9.03 | 9.03
Ccal_0 | 0 | mg/L | 0.0057735 |  |  | 0 | 0 | 0.0057735 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ccal | 8.9792 | mg/L | 0.031 |  |  | 8.9792 | 8.9792 | 8.9792 | 9.0102 | 8.9792 | 8.9792 | 8.9792 | 8.9792 | 8.9792 | 8.9792
Cread_meas | 11.95 | mg/L | 0.068705 |  |  | 11.95 | 11.95 | 11.95 | 11.95 | 12.0187 | 11.95 | 11.95 | 11.95 | 11.95 | 11.95
γ | 1.0016 | - | 0.00045357 |  |  | 1.0016 | 1.0016 | 1.0016 | 1.0016 | 1.0016 | 1.002 | 1.0016 | 1.0016 | 1.0016 | 1.0016
ΔCstir | 0 | mg/L | 0.0032511 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.0032511 | 0 | 0 | 0
ΔCΔt | 0 | mg/L | 0.0015132 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0015132 | 0 | 0
ΔCdrift | 0 | mg/L | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
ΔClinearity | 0 | mg/L | 0.015174 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.015174
Cresult | 11.8642 | mg/L |  |  |  | 11.866 | 11.8578 | 11.8623 | 11.9051 | 11.9324 | 11.8588 | 11.8674 | 11.8657 | 11.8642 | 11.8793
uCresult | 0.081525 | mg/L |  |  |  | 0.0018217 | -0.0063743 | -0.001855 | 0.04096 | 0.068211 | -0.0053704 | 0.0032278 | 0.0015023 | 0 | 0.015065
UCresult | 0.16305 | mg/L, k=2 |  |  |  | 3.3185e-06 | 4.0632e-05 | 3.4409e-06 | 0.0016777 | 0.0046528 | 2.8841e-05 | 1.0419e-05 | 2.2569e-06 | 0 | 0.00022695
UCresult rel | 1.3743 | %, k=2 | [12 empty] | 0.0066464
| En (vs WTW) = | -0.1509 |  |  |  | 0.04993% | 0.6113% | 0.05177% | 25.24% | 70% | 0.4339% | 0.1568% | 0.03396% | 0% | 3.415%
100%

## Table S7
Table S7. Measurement uncertainty using the ISO GUM method with Kragten approach
Data from Lake Ilmatsalu 2
WTW
Experimental data for calculating different uncertainty contributions:
u(Ccal_rep) | 0.019833 | mg/L | Pooled standard deviation of the readings at stirring rate 120 rpm
u(Ccal_0_rep) | 0.019833 | mg/L | The largest pooled standard deviation of the readings
u(Cmeas_Rep) | 0.051188 | mg/L | Reproducibility: pooled standard deviation of differences between reading and reference value
u(Ccal) | 0.031 | mg/L | From Ref 3
b1stir | 0.014614 | %/% | The slope of stirring graph (% of the reading vs % of the stirring rate from calibration stirring rate)
b1Δt | -0.077609 | %/°C | The slope of temperature compensation graph (temperature °C vs relative difference from reference temperature in %)
b1drift | 0.0037617 | mg/L | The largest slope of the drift graph (number of days passed from calibration vs reading difference from the reference value mg/L)
s(b1drift) | 0.00089586 | mg/L | Standard deviation of the slope of the drift graph
Input quantities
| Value | Unit | Comment
Calibration
Cread_cal_0 | 0 | mg/L | Sensor reading right after calibration/checking zero value | [7 empty] | Constants for salinity calculation (Ref 6):
Cread_cal | 8.98 | mg/L | Sensor reading right after calibration | [7 empty] | a0 | 0.008
Ccal_0 | 0 | mg/L | Calibration environment DO concentration at oxygen-free conditions, here always 0 | [7 empty] | a1 | -0.1692
Ccal | 8.9792 | mg/L | Calibration envirnment DO concentration at saturation, calculated by the Ref 4 | [7 empty] | a2 | 25.3851 |  | C(KCl at 25°C) *
tcal_ref | 19.909 | °C | Temperature of calibration solution measured with reference thermometer | [7 empty] | a3 | 14.0941 |  | 53.9038 | mS/cm
Pcal_ref | 99915 | Pa | Air pressure at the time of calibration measured with reference barometer | [7 empty] | a4 | -7.0261 |  | *Notes:
tcal_sensor | 19.9 | °C | Temperature of calibration solution measured with DO-meter | [7 empty] | a5 | 2.7081 |  | Ref 5: | conductivity of SSW at SP=35, t_68=15, p=0; 42.9140 mS/cm
Pcal_sensor | 100100 | Pa | Air pressure at the time of calibration measured with DO-meter | [7 empty] | b0 | 0.0005 |  | Ref 7: | factor for calculating the conductivity from 15 °C to the 25 °C is found as 1.25608904610271
[11 empty] | b1 | -0.0056
Measurement | [10 empty] | b2 | -0.0066
Cread_meas | 11.92 | mg/L | Reading of the measured sample | [7 empty] | b3 | -0.0375
k | 0.558 | mS/cm | conductivity of the sample, used for calculating the activity coefficient | [7 empty] | b4 | 0.0636
Salinity | 0.26449 | PSU | Salinity, calculated from measured conductivity (recalculated to 25 C), using Refs 5,6 and 7 | [7 empty] | b5 | -0.0144
tsample | 18.9 | °C | Sample temperature, measured with DO-meter | [7 empty] | X | 400 | *Rt
Δvstir | 50 | % | estimated flow rate compared to the flow rate of calibration | [7 empty] | Y | 100 | *Rt
Δdays | 0 |  | The number of days passed from the last calibration
Model components |  |  |  |  | Uncertainty components
Denotion | Value | Unit | uc | Relative | u1 |  |  | u2
Cread_cal_0 | 0 | mg/L | 0.020042 |  | 0.0028868 | u(Ccal_res) | ±half of the last digit | 0.019833 | u(Ccal_rep) | repeatabilty of the reading
Cread_cal | 8.98 | mg/L | 0.020042 | 0.2232% | 0.0028868 | u(Ccal_0_res) | ±half of the last digit | 0.019833 | u(Ccal_0_rep) | repeatabilty of the reading
Ccal_0 | 0 | mg/L | 0.0057735 |  | 0.0057735 | u(Ccal_0) | estimation
Ccal | 8.9792 | mg/L | 0.031 | 0.3452% | 0.031 | u(Ccal) | Estimated as in Ref 3
Cread_meas | 11.92 | mg/L | 0.051269 | 0.4301% | 0.0028868 | u(Cmeas_res) | ±half of the last digit | 0.051188 | u(Cmeas_Rep) | reproducibility of the reading
γ | 1.0016 | - | 0.00045462 | 0.04539% | 0.00045462 | ± half of the activity coefficient difference from 1
ΔCstir | 0 | mg/L | 0.050288 |  | 0.050288 | ±Δvstir *b1stir*Cread_meas/100%
ΔCΔt | 0 | mg/L | 0.0053411 |  | 0.0053411 | ±Δt*b1Δt/100%*Cread_meas
ΔCdrift | 0 | mg/L | 0 |  | 0 | ±Δdays*(b1drift+s(b1drift))
Denotion | Value | Unit | uc |  |  | Cread_cal_0 | Cread_cal | Ccal_0 | Ccal | Cread_meas | γ | ΔCstir | ΔCΔt | ΔCdrift
Cread_cal_0 | 0 | mg/L | 0.020042 |  |  | 0.020042 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cread_cal | 8.98 | mg/L | 0.020042 |  |  | 8.98 | 9 | 8.98 | 8.98 | 8.98 | 8.98 | 8.98 | 8.98 | 8.98
Ccal_0 | 0 | mg/L | 0.0057735 |  |  | 0 | 0 | 0.0057735 | 0 | 0 | 0 | 0 | 0 | 0
Ccal | 8.9792 | mg/L | 0.031 |  |  | 8.9792 | 8.9792 | 8.9792 | 9.0102 | 8.9792 | 8.9792 | 8.9792 | 8.9792 | 8.9792
Cread_meas | 11.92 | mg/L | 0.051269 |  |  | 11.92 | 11.92 | 11.92 | 11.92 | 11.9713 | 11.92 | 11.92 | 11.92 | 11.92
γ | 1.0016 | - | 0.00045462 |  |  | 1.0016 | 1.0016 | 1.0016 | 1.0016 | 1.0016 | 1.002 | 1.0016 | 1.0016 | 1.0016
ΔCstir | 0 | mg/L | 0.050288 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.050288 | 0 | 0
ΔCΔt | 0 | mg/L | 0.0053411 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0053411 | 0
ΔCdrift | 0 | mg/L | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cresult | 11.9003 | mg/L |  |  |  | 11.9068 | 11.8738 | 11.8984 | 11.9413 | 11.9514 | 11.8949 | 11.9505 | 11.9056 | 11.9003
uCresult | 0.087375 | mg/L |  |  |  | 0.0065337 | -0.0265 | -0.0018782 | 0.041085 | 0.051184 | -0.0053991 | 0.050204 | 0.0053322 | 0
UCresult | 0.17475 | mg/L, k=2 |  |  |  | 4.2689e-05 | 0.00070225 | 3.5275e-06 | 0.0016879 | 0.0026198 | 2.915e-05 | 0.0025205 | 2.8433e-05 | 0
UCresult rel | 1.4684 | %, k=2 | [11 empty] | 0.0076343
[6 empty] | 0.5592% | 9.199% | 0.04621% | 22.11% | 34.32% | 0.3818% | 33.02% | 0.3724% | 0%
100%

## Table S8
Table S8. Measurement uncertainty using the ISO GUM method with Kragten approach
Data from Jordan spring
HACH
Experimental data for calculating different uncertainty contributions:
u(Ccal_rep) | 0.0039025 | mg/L | Pooled standard deviation of the readings at stirring rate 120 rpm
u(Ccal_0_rep) | 0.004913 | mg/L | The largest pooled standard deviation of the readings
u(Cmeas_Rep) | 0.068644 | mg/L | Reproducibility: pooled standard deviation of differences between reading and reference value
u(Ccal) | 0.031 | mg/L | From Ref 3
b1stir | -0.00094245 | %/% | The slope of stirring graph (% of the reading vs % of the stirring rate from calibration stirring rate)
b1Δt | 0.024369 | %/°C | The slope of temperature compensation graph (temperature °C vs relative difference from reference temperature in %)
b1drift | -0.0078995 | mg/L | The slope of the drift graph (number of days passed from calibration vs reading difference from the reference value mg/L)
s(b1drift) | 0.00084931 | mg/L | Standard deviation of the slope of the drift graph
b1linearity | 0.027665 | - | The slope of the linearity graph starting at 11 mg/L (electrochemical reading mg/L vs readings difference mg/L)
Input quantities
| Value | Unit | Comment
Calibration
Cread_cal_0 | 0 | mg/L | Sensor reading right after calibration/checking zero value
Cread_cal | 9.03 | mg/L | Sensor reading right after calibration | [7 empty] | Constants for salinity calculation (Ref 6):
Ccal_0 | 0 | mg/L | Calibration environment DO concentration at oxygen-free conditions, here always 0 | [7 empty] | a0 | 0.008
Ccal | 8.9792 | mg/L | Calibration envirnment DO concentration at saturation, calculated by the Ref 4 | [7 empty] | a1 | -0.1692 |  | C(KCl at 25°C) *
tcal_ref | 19.909 | °C | Temperature of calibration solution measured with reference thermometer | [7 empty] | a2 | 25.3851 |  | 53.9038 | mS/cm
Pcal_ref | 99915 | Pa | Air pressure at the time of calibration measured with reference barometer | [7 empty] | a3 | 14.0941 |  | *Notes:
tcal_sensor | 20.1 | °C | Temperature of calibration solution measured with DO-meter | [7 empty] | a4 | -7.0261 |  | Ref 5: | conductivity of SSW at SP=35, t_68=15, p=0; 42.9140 mS/cm
Pcal_sensor | 100100 | Pa | Air pressure at the time of calibration measured with DO-meter | [7 empty] | a5 | 2.7081 |  | Ref 7: | factor for calculating the conductivity from 15 °C to the 25 °C is found as 1.25608904610271
[11 empty] | b0 | 0.0005
Measurement | [10 empty] | b1 | -0.0056
Cread_meas | 7.71 | mg/L | Reading of the measured sample | [7 empty] | b2 | -0.0066
k | 0.896 | mS/cm | conductivity of the sample, used for calculating the activity coefficient | [7 empty] | b3 | -0.0375
Salinity | 0.43199 | PSU | Salinity, calculated from measured conductivity (recalculated to 25 C), using Refs 5, 6 and 7 | [7 empty] | b4 | 0.0636
tsample | 7.1 | °C | Sample temperature, measured with DO-meter | [7 empty] | b5 | -0.0144
Δvstir | 50 | % | estimated flow rate compared to the flow rate of calibration | [7 empty] | X | 400 | *Rt
Δdays | 2 |  | The number of days passed from the last calibration | [7 empty] | Y | 100 | *Rt
Model components |  |  |  |  | Uncertainty components
Denotion | Value | Unit | uc | Relative | u1 |  |  | u2
Cread_cal_0 | 0 | mg/L | 0.0056984 |  | 0.0028868 | u(Ccal_res) | ±half of the last digit | 0.004913 | u(Ccal_rep) | repeatabilty of the reading
Cread_cal | 9.03 | mg/L | 0.0048542 | 0.05376% | 0.0028868 | u(Ccal_0_res) | ±half of the last digit | 0.0039025 | u(Ccal_0_rep) | repeatabilty of the reading
Ccal_0 | 0 | mg/L | 0.0057735 |  | 0.0057735 | u(Ccal_0) | estimation
Ccal | 8.9792 | mg/L | 0.031 | 0.3452% | 0.031 | u(Ccal) | Estimated as in Ref 3
Cread_meas | 7.71 | mg/L | 0.068705 | 0.8911% | 0.0028868 | u(Cmeas_res) | ±half of the last digit | 0.068644 | u(Cmeas_Rep) | Reproducibility of the reading
γ | 1.0028 | - | 0.00081888 | 0.08166% | 0.00081888 | ± half of the activity coefficient difference from 1
ΔCstir | 0 | mg/L | 0.0020976 |  | 0.0020976 | ±Δvstir *b1stir*Cread_meas/100%
ΔCΔt | 0 | mg/L | 0.014102 |  | 0.014102 | ±Δt*b1Δt/100%*Cread_meas
ΔCdrift | 0 | mg/L | 0.010102 |  | 0.010102 | ±Δdays*(b1drift+s(b1drift))
ΔClinearity | 0 | mg/L | 0 |  | 0 | ±ΔC from 11 mg/L*b1linearity (if Cread_meas>11)
Denotion | Value | Unit | uc |  |  | Cread_cal_0 | Cread_cal | Ccal_0 | Ccal | Cread_meas | γ | ΔCstir | ΔCΔt | ΔCdrift | ΔClinearity
Cread_cal_0 | 0 | mg/L | 0.0056984 |  |  | 0.0056984 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cread_cal | 9.03 | mg/L | 0.0048542 |  |  | 9.03 | 9.0349 | 9.03 | 9.03 | 9.03 | 9.03 | 9.03 | 9.03 | 9.03 | 9.03
Ccal_0 | 0 | mg/L | 0.0057735 |  |  | 0 | 0 | 0.0057735 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ccal | 8.9792 | mg/L | 0.031 |  |  | 8.9792 | 8.9792 | 8.9792 | 9.0102 | 8.9792 | 8.9792 | 8.9792 | 8.9792 | 8.9792 | 8.9792
Cread_meas | 7.71 | mg/L | 0.068705 |  |  | 7.71 | 7.71 | 7.71 | 7.71 | 7.7787 | 7.71 | 7.71 | 7.71 | 7.71 | 7.71
γ | 1.0028 | - | 0.00081888 |  |  | 1.0028 | 1.0028 | 1.0028 | 1.0028 | 1.0028 | 1.0037 | 1.0028 | 1.0028 | 1.0028 | 1.0028
ΔCstir | 0 | mg/L | 0.0020976 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.0020976 | 0 | 0 | 0
ΔCΔt | 0 | mg/L | 0.014102 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.014102 | 0 | 0
ΔCdrift | 0 | mg/L | 0.010102 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.010102 | 0
ΔClinearity | 0 | mg/L | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cresult | 7.645 | mg/L |  |  |  | 7.6441 | 7.6409 | 7.6458 | 7.6714 | 7.7131 | 7.6387 | 7.6471 | 7.659 | 7.655 | 7.645
uCresult | 0.075466 | mg/L |  |  |  | -0.00084252 | -0.0041075 | 0.00085791 | 0.026394 | 0.068125 | -0.0062375 | 0.0020799 | 0.013983 | 0.010017 | 0
UCresult | 0.15093 | mg/L, k=2 |  |  |  | 7.0983e-07 | 1.6871e-05 | 7.3601e-07 | 0.00069662 | 0.004641 | 3.8906e-05 | 4.326e-06 | 0.00019552 | 0.00010034 | 0
UCresult rel | 1.9743 | %, k=2 | [12 empty] | 0.0056951
| En (vs WTW) = | -0.36882 |  |  |  | 0.01246% | 0.2962% | 0.01292% | 12.23% | 81.49% | 0.6832% | 0.07596% | 3.433% | 1.762% | 0%
100%

## Table S9
Table S9. Measurement uncertainty using the ISO GUM method with Kragten approach
Data from Jordan spring
WTW
Experimental data for calculating different uncertainty contributions:
u(Ccal_rep) | 0.019833 | mg/L | Pooled standard deviation of the readings at stirring rate 120 rpm
u(Ccal_0_rep) | 0.019833 | mg/L | The largest pooled standard deviation of the readings
u(Cmeas_Rep) | 0.051188 | mg/L | Reproducibility: pooled standard deviation of differences between reading and reference value
u(Ccal) | 0.031 | mg/L | From Ref 3
b1stir | 0.014614 | %/% | The slope of stirring graph (% of the reading vs % of the stirring rate from calibration stirring rate)
b1Δt | -0.077609 | %/°C | The slope of temperature compensation graph (temperature °C vs relative difference from reference temperature in %)
b1drift | 0.0037617 | mg/L | The largest slope of the drift graph (number of days passed from calibration vs reading difference from the reference value mg/L)
s(b1drift) | 0.00089586 | mg/L | Standard deviation of the slope of the drift graph
Input quantities
| Value | Unit | Comment
Calibration
Cread_cal_0 | 0 | mg/L | Sensor reading right after calibration/checking zero value | [7 empty] | Constants for salinity calculation (Ref 6):
Cread_cal | 8.98 | mg/L | Sensor reading right after calibration | [7 empty] | a0 | 0.008
Ccal_0 | 0 | mg/L | Calibration environment DO concentration at oxygen-free conditions, here always 0 | [7 empty] | a1 | -0.1692
Ccal | 8.9792 | mg/L | Calibration envirnment DO concentration at saturation, calculated by the Ref 4 | [7 empty] | a2 | 25.3851 |  | C(KCl at 25°C) *
tcal_ref | 19.909 | °C | Temperature of calibration solution measured with reference thermometer | [7 empty] | a3 | 14.0941 |  | 53.9038 | mS/cm
Pcal_ref | 99915 | Pa | Air pressure at the time of calibration measured with reference barometer | [7 empty] | a4 | -7.0261 |  | *Notes:
tcal_sensor | 19.9 | °C | Temperature of calibration solution measured with DO-meter | [7 empty] | a5 | 2.7081 |  | Ref 5: | conductivity of SSW at SP=35, t_68=15, p=0; 42.9140 mS/cm
Pcal_sensor | 100100 | Pa | Air pressure at the time of calibration measured with DO-meter | [7 empty] | b0 | 0.0005 |  | Ref 7: | factor for calculating the conductivity from 15 °C to the 25 °C is found as 1.25608904610271
[11 empty] | b1 | -0.0056
Measurement | [10 empty] | b2 | -0.0066
Cread_meas | 7.75 | mg/L | Reading of the measured sample | [7 empty] | b3 | -0.0375
k | 0.896 | mS/cm | conductivity of the sample, used for calculating the activity coefficient | [7 empty] | b4 | 0.0636
Salinity | 0.43199 | PSU | Salinity, calculated from measured conductivity (recalculated to 25 C), using Refs 5, 6 and 7 | [7 empty] | b5 | -0.0144
tsample | 6.9 | °C | Sample temperature, measured with DO-meter | [7 empty] | X | 400 | *Rt
Δvstir | 50 | % | estimated flow rate compared to the flow rate of calibration | [7 empty] | Y | 100 | *Rt
Δdays | 2 |  | The number of days passed from the last calibration
Model components |  |  |  |  | Uncertainty components
Denotion | Value | Unit | uc | Relative | u1 |  |  | u2
Cread_cal_0 | 0 | mg/L | 0.020042 |  | 0.0028868 | u(Ccal_res) | ±half of the last digit | 0.019833 | u(Ccal_rep) | repeatabilty of the reading
Cread_cal | 8.98 | mg/L | 0.020042 | 0.2232% | 0.0028868 | u(Ccal_0_res) | ±half of the last digit | 0.019833 | u(Ccal_0_rep) | repeatabilty of the reading
Ccal_0 | 0 | mg/L | 0.0057735 |  | 0.0057735 | u(Ccal_0) | estimation
Ccal | 8.9792 | mg/L | 0.031 | 0.3452% | 0.031 | u(Ccal) | Estimated as in Ref 3
Cread_meas | 7.75 | mg/L | 0.051269 | 0.6615% | 0.0028868 | u(Cmeas_res) | ±half of the last digit | 0.051188 | u(Cmeas_Rep) | reproducibility of the reading
γ | 1.0028 | - | 0.00082032 | 0.0818% | 0.00082032 | ± half of the activity coefficient difference from 1
ΔCstir | 0 | mg/L | 0.032695 |  | 0.032695 | ±Δvstir *b1stir*Cread_meas/100%
ΔCΔt | 0 | mg/L | 0.045144 |  | 0.045144 | ±Δt*b1Δt/100%*Cread_meas
ΔCdrift | 0 | mg/L | 0.0053781 |  | 0.0053781 | ±Δdays*(b1drift+s(b1drift))
Denotion | Value | Unit | uc |  |  | Cread_cal_0 | Cread_cal | Ccal_0 | Ccal | Cread_meas | γ | ΔCstir | ΔCΔt | ΔCdrift
Cread_cal_0 | 0 | mg/L | 0.020042 |  |  | 0.020042 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cread_cal | 8.98 | mg/L | 0.020042 |  |  | 8.98 | 9 | 8.98 | 8.98 | 8.98 | 8.98 | 8.98 | 8.98 | 8.98
Ccal_0 | 0 | mg/L | 0.0057735 |  |  | 0 | 0 | 0.0057735 | 0 | 0 | 0 | 0 | 0 | 0
Ccal | 8.9792 | mg/L | 0.031 |  |  | 8.9792 | 8.9792 | 8.9792 | 9.0102 | 8.9792 | 8.9792 | 8.9792 | 8.9792 | 8.9792
Cread_meas | 7.75 | mg/L | 0.051269 |  |  | 7.75 | 7.75 | 7.75 | 7.75 | 7.8013 | 7.75 | 7.75 | 7.75 | 7.75
γ | 1.0028 | - | 0.00082032 |  |  | 1.0028 | 1.0028 | 1.0028 | 1.0028 | 1.0028 | 1.0037 | 1.0028 | 1.0028 | 1.0028
ΔCstir | 0 | mg/L | 0.032695 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.032695 | 0 | 0
ΔCΔt | 0 | mg/L | 0.045144 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.045144 | 0
ΔCdrift | 0 | mg/L | 0.0053781 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0053781
Cresult | 7.7274 | mg/L |  |  |  | 7.7246 | 7.7102 | 7.7282 | 7.7541 | 7.7785 | 7.7211 | 7.76 | 7.7724 | 7.7328
uCresult | 0.082384 | mg/L |  |  |  | -0.0028001 | -0.017208 | 0.00080492 | 0.026678 | 0.05112 | -0.0063158 | 0.0326 | 0.045012 | 0.0053624
UCresult | 0.16477 | mg/L, k=2 |  |  |  | 7.8408e-06 | 0.0002961 | 6.479e-07 | 0.00071172 | 0.0026132 | 3.9889e-05 | 0.0010628 | 0.0020261 | 2.8756e-05
UCresult rel | 2.1322 | %, k=2 | [11 empty] | 0.006787
[6 empty] | 0.1155% | 4.363% | 0.009546% | 10.49% | 38.5% | 0.5877% | 15.66% | 29.85% | 0.4237%
100%

## Table S10
Table S10. Measurement uncertainty using the ISO GUM method with Kragten approach
Data from Pärnu coast
HACH
Experimental data for calculating different uncertainty contributions:
u(Ccal_rep) | 0.0039025 | mg/L | Pooled standard deviation of the readings at stirring rate 120 rpm
u(Ccal_0_rep) | 0.004913 | mg/L | The largest pooled standard deviation of the readings
u(Cmeas_Rep) | 0.068644 | mg/L | Reproducibility: pooled standard deviation of differences between reading and reference value
u(Ccal) | 0.031 | mg/L | From Ref 3
b1stir | -0.00094245 | %/% | The slope of stirring graph (% of the reading vs % of the stirring rate from calibration stirring rate)
b1Δt | 0.024369 | %/°C | The slope of temperature compensation graph (temperature °C vs relative difference from reference temperature in %)
b1drift | -0.0078995 | mg/L | The slope of the drift graph (number of days passed from calibration vs reading difference from the reference value mg/L)
s(b1drift) | 0.00084931 | mg/L | Standard deviation of the slope of the drift graph
b1linearity | 0.027665 | - | The slope of the linearity graph starting at 11 mg/L (electrochemical reading mg/L vs readings difference mg/L)
Input quantities
| Value | Unit | Comment
Calibration
Cread_cal_0 | 0 | mg/L | Sensor reading right after calibration/checking zero value
Cread_cal | 9.03 | mg/L | Sensor reading right after calibration | [7 empty] | Constants for salinity calculation (Ref 6):
Ccal_0 | 0 | mg/L | Calibration environment DO concentration at oxygen-free conditions, here always 0 | [7 empty] | a0 | 0.008
Ccal | 8.9792 | mg/L | Calibration envirnment DO concentration at saturation, calculated by the Ref 4 | [7 empty] | a1 | -0.1692 |  | C(KCl at 25°C) *
tcal_ref | 19.909 | °C | Temperature of calibration solution measured with reference thermometer | [7 empty] | a2 | 25.3851 |  | 53.9038 | mS/cm
Pcal_ref | 99915 | Pa | Air pressure at the time of calibration measured with reference barometer | [7 empty] | a3 | 14.0941 |  | *Notes:
tcal_sensor | 20.1 | °C | Temperature of calibration solution measured with DO-meter | [7 empty] | a4 | -7.0261 |  | Ref 5: | conductivity of SSW at SP=35, t_68=15, p=0; 42.9140 mS/cm
Pcal_sensor | 100100 | Pa | Air pressure at the time of calibration measured with DO-meter | [7 empty] | a5 | 2.7081 |  | Ref 7: | factor for calculating the conductivity from 15 °C to the 25 °C is found as 1.25608904610271
[11 empty] | b0 | 0.0005
Measurement | [10 empty] | b1 | -0.0056
Cread_meas | 11.12 | mg/L | Reading of the measured sample | [7 empty] | b2 | -0.0066
k | 3.1 | mS/cm | conductivity of the sample, used for calculating the activity coefficient | [7 empty] | b3 | -0.0375
Salinity | 1.5871 | PSU | Salinity, calculated from measured conductivity (recalculated to 25 C), using Refs 5, 6 and 7 | [7 empty] | b4 | 0.0636
tsample | 18.4 | °C | Sample temperature, measured with DO-meter | [7 empty] | b5 | -0.0144
Δvstir | 50 | % | estimated flow rate compared to the flow rate of calibration | [7 empty] | X | 400 | *Rt
Δdays | 3 |  | The number of days passed from the last calibration | [7 empty] | Y | 100 | *Rt
Model components |  |  |  |  | Uncertainty components
Denotion | Value | Unit | uc | Relative | u1 |  |  | u2
Cread_cal_0 | 0 | mg/L | 0.0056984 |  | 0.0028868 | u(Ccal_res) | ±half of the last digit | 0.004913 | u(Ccal_rep) | repeatabilty of the reading
Cread_cal | 9.03 | mg/L | 0.0048542 | 0.05376% | 0.0028868 | u(Ccal_0_res) | ±half of the last digit | 0.0039025 | u(Ccal_0_rep) | repeatabilty of the reading
Ccal_0 | 0 | mg/L | 0.0057735 |  | 0.0057735 | u(Ccal_0) | estimation
Ccal | 8.9792 | mg/L | 0.031 | 0.3452% | 0.031 | u(Ccal) | Estimated as in Ref 3
Cread_meas | 11.12 | mg/L | 0.068705 | 0.6178% | 0.0028868 | u(Cmeas_res) | ±half of the last digit | 0.068644 | u(Cmeas_Rep) | Reproducibility of the reading
γ | 1.0095 | - | 0.0027494 | 0.2723% | 0.0027494 | ± half of the activity coefficient difference from 1
ΔCstir | 0 | mg/L | 0.0030253 |  | 0.0030253 | ±Δvstir *b1stir*Cread_meas/100%
ΔCΔt | 0 | mg/L | 0.0026597 |  | 0.0026597 | ±Δt*b1Δt/100%*Cread_meas
ΔCdrift | 0 | mg/L | 0.015153 |  | 0.015153 | ±Δdays*(b1drift+s(b1drift))
ΔClinearity | 0 | mg/L | 0.0019167 |  | 0.0019167 | ±ΔC from 11 mg/L*b1linearity (if Cread_meas>11)
Denotion | Value | Unit | uc |  |  | Cread_cal_0 | Cread_cal | Ccal_0 | Ccal | Cread_meas | γ | ΔCstir | ΔCΔt | ΔCdrift | ΔClinearity
Cread_cal_0 | 0 | mg/L | 0.0056984 |  |  | 0.0056984 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cread_cal | 9.03 | mg/L | 0.0048542 |  |  | 9.03 | 9.0349 | 9.03 | 9.03 | 9.03 | 9.03 | 9.03 | 9.03 | 9.03 | 9.03
Ccal_0 | 0 | mg/L | 0.0057735 |  |  | 0 | 0 | 0.0057735 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ccal | 8.9792 | mg/L | 0.031 |  |  | 8.9792 | 8.9792 | 8.9792 | 9.0102 | 8.9792 | 8.9792 | 8.9792 | 8.9792 | 8.9792 | 8.9792
Cread_meas | 11.12 | mg/L | 0.068705 |  |  | 11.12 | 11.12 | 11.12 | 11.12 | 11.1887 | 11.12 | 11.12 | 11.12 | 11.12 | 11.12
γ | 1.0095 | - | 0.0027494 |  |  | 1.0095 | 1.0095 | 1.0095 | 1.0095 | 1.0095 | 1.0123 | 1.0095 | 1.0095 | 1.0095 | 1.0095
ΔCstir | 0 | mg/L | 0.0030253 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.0030253 | 0 | 0 | 0
ΔCΔt | 0 | mg/L | 0.0026597 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0026597 | 0 | 0
ΔCdrift | 0 | mg/L | 0.015153 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.015153 | 0
ΔClinearity | 0 | mg/L | 0.0019167 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0019167
Cresult | 10.9532 | mg/L |  |  |  | 10.9544 | 10.9473 | 10.9519 | 10.991 | 11.0209 | 10.9234 | 10.9562 | 10.9558 | 10.9681 | 10.9551
uCresult | 0.084703 | mg/L |  |  |  | 0.0012464 | -0.0058849 | -0.0012692 | 0.037815 | 0.067674 | -0.02975 | 0.0029799 | 0.0026198 | 0.014926 | 0.0018879
UCresult | 0.16941 | mg/L, k=2 |  |  |  | 1.5536e-06 | 3.4632e-05 | 1.6109e-06 | 0.00143 | 0.0045798 | 0.00088504 | 8.88e-06 | 6.8631e-06 | 0.00022279 | 3.5643e-06
UCresult rel | 1.5466 | %, k=2 | [12 empty] | 0.0071746
| En (vs WTW) = | -0.40888 |  |  |  | 0.02165% | 0.4827% | 0.02245% | 19.93% | 63.83% | 12.34% | 0.1238% | 0.09566% | 3.105% | 0.04968%
100%

## Table S11
Table S11. Measurement uncertainty using the ISO GUM method with Kragten approach
Data from Pärnu coast
WTW
Experimental data for calculating different uncertainty contributions:
u(Ccal_rep) | 0.019833 | mg/L | Pooled standard deviation of the readings at stirring rate 120 rpm
u(Ccal_0_rep) | 0.019833 | mg/L | The largest pooled standard deviation of the readings
u(Cmeas_Rep) | 0.051188 | mg/L | Reproducibility: pooled standard deviation of differences between reading and reference value
u(Ccal) | 0.031 | mg/L | From Ref 3
b1stir | 0.014614 | %/% | The slope of stirring graph (% of the reading vs % of the stirring rate from calibration stirring rate)
b1Δt | -0.077609 | %/°C | The slope of temperature compensation graph (temperature °C vs relative difference from reference temperature in %)
b1drift | 0.0037617 | mg/L | The largest slope of the drift graph (number of days passed from calibration vs reading difference from the reference value mg/L)
s(b1drift) | 0.00089586 | mg/L | Standard deviation of the slope of the drift graph
Input quantities
| Value | Unit | Comment
Calibration
Cread_cal_0 | 0 | mg/L | Sensor reading right after calibration/checking zero value | [7 empty] | Constants for salinity calculation (Ref 6):
Cread_cal | 8.98 | mg/L | Sensor reading right after calibration | [7 empty] | a0 | 0.008
Ccal_0 | 0 | mg/L | Calibration environment DO concentration at oxygen-free conditions, here always 0 | [7 empty] | a1 | -0.1692
Ccal | 8.9792 | mg/L | Calibration envirnment DO concentration at saturation, calculated by the Ref 4 | [7 empty] | a2 | 25.3851 |  | C(KCl at 25°C) *
tcal_ref | 19.909 | °C | Temperature of calibration solution measured with reference thermometer | [7 empty] | a3 | 14.0941 |  | 53.9038 | mS/cm
Pcal_ref | 99915 | Pa | Air pressure at the time of calibration measured with reference barometer | [7 empty] | a4 | -7.0261 |  | *Notes:
tcal_sensor | 19.9 | °C | Temperature of calibration solution measured with DO-meter | [7 empty] | a5 | 2.7081 |  | Ref 5: | conductivity of SSW at SP=35, t_68=15, p=0; 42.9140 mS/cm
Pcal_sensor | 100100 | Pa | Air pressure at the time of calibration measured with DO-meter | [7 empty] | b0 | 0.0005 |  | Ref 7: | factor for calculating the conductivity from 15 °C to the 25 °C is found as 1.25608904610271
[11 empty] | b1 | -0.0056
Measurement | [10 empty] | b2 | -0.0066
Cread_meas | 11.16 | mg/L | Reading of the measured sample | [7 empty] | b3 | -0.0375
k | 3.1 | mS/cm | conductivity of the sample, used for calculating the activity coefficient | [7 empty] | b4 | 0.0636
Salinity | 1.5871 | PSU | Salinity, calculated from measured conductivity (recalculated to 25 C), using Refs 5, 6 and 7 | [7 empty] | b5 | -0.0144
tsample | 18.1 | °C | Sample temperature, measured with DO-meter | [7 empty] | X | 400 | *Rt
Δvstir | 50 | % | estimated flow rate compared to the flow rate of calibration | [7 empty] | Y | 100 | *Rt
Δdays | 3 |  | The number of days passed from the last calibration
Model components |  |  |  |  | Uncertainty components
Denotion | Value | Unit | uc | Relative | u1 |  |  | u2
Cread_cal_0 | 0 | mg/L | 0.020042 |  | 0.0028868 | u(Ccal_res) | ±half of the last digit | 0.019833 | u(Ccal_rep) | repeatabilty of the reading
Cread_cal | 8.98 | mg/L | 0.020042 | 0.2232% | 0.0028868 | u(Ccal_0_res) | ±half of the last digit | 0.019833 | u(Ccal_0_rep) | repeatabilty of the reading
Ccal_0 | 0 | mg/L | 0.0057735 |  | 0.0057735 | u(Ccal_0) | estimation
Ccal | 8.9792 | mg/L | 0.031 | 0.3452% | 0.031 | u(Ccal) | Estimated as in Ref 3
Cread_meas | 11.16 | mg/L | 0.051269 | 0.4594% | 0.0028868 | u(Cmeas_res) | ±half of the last digit | 0.051188 | u(Cmeas_Rep) | reproducibility of the reading
γ | 1.0095 | - | 0.0027558 | 0.273% | 0.0027558 | ± half of the activity coefficient difference from 1
ΔCstir | 0 | mg/L | 0.047082 |  | 0.047082 | ±Δvstir *b1stir*Cread_meas/100%
ΔCΔt | 0 | mg/L | 0.009001 |  | 0.009001 | ±Δt*b1Δt/100%*Cread_meas
ΔCdrift | 0 | mg/L | 0.0080672 |  | 0.0080672 | ±Δdays*(b1drift+s(b1drift))
Denotion | Value | Unit | uc |  |  | Cread_cal_0 | Cread_cal | Ccal_0 | Ccal | Cread_meas | γ | ΔCstir | ΔCΔt | ΔCdrift
Cread_cal_0 | 0 | mg/L | 0.020042 |  |  | 0.020042 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cread_cal | 8.98 | mg/L | 0.020042 |  |  | 8.98 | 9 | 8.98 | 8.98 | 8.98 | 8.98 | 8.98 | 8.98 | 8.98
Ccal_0 | 0 | mg/L | 0.0057735 |  |  | 0 | 0 | 0.0057735 | 0 | 0 | 0 | 0 | 0 | 0
Ccal | 8.9792 | mg/L | 0.031 |  |  | 8.9792 | 8.9792 | 8.9792 | 9.0102 | 8.9792 | 8.9792 | 8.9792 | 8.9792 | 8.9792
Cread_meas | 11.16 | mg/L | 0.051269 |  |  | 11.16 | 11.16 | 11.16 | 11.16 | 11.2113 | 11.16 | 11.16 | 11.16 | 11.16
γ | 1.0095 | - | 0.0027558 |  |  | 1.0095 | 1.0095 | 1.0095 | 1.0095 | 1.0095 | 1.0123 | 1.0095 | 1.0095 | 1.0095
ΔCstir | 0 | mg/L | 0.047082 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.047082 | 0 | 0
ΔCΔt | 0 | mg/L | 0.009001 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.009001 | 0
ΔCdrift | 0 | mg/L | 0.0080672 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0080672
Cresult | 11.0535 | mg/L |  |  |  | 11.0582 | 11.0289 | 11.0522 | 11.0917 | 11.1043 | 11.0234 | 11.1002 | 11.0625 | 11.0615
uCresult | 0.088812 | mg/L |  |  |  | 0.0046398 | -0.024614 | -0.0013337 | 0.038161 | 0.05078 | -0.030092 | 0.046632 | 0.0089151 | 0.0079903
UCresult | 0.17762 | mg/L, k=2 |  |  |  | 2.1528e-05 | 0.00060587 | 1.7789e-06 | 0.0014563 | 0.0025786 | 0.0009055 | 0.0021746 | 7.948e-05 | 6.3844e-05
UCresult rel | 1.6069 | %, k=2 | [11 empty] | 0.0078875
[6 empty] | 0.2729% | 7.681% | 0.02255% | 18.46% | 32.69% | 11.48% | 27.57% | 1.008% | 0.8094%
100%

## Table S12
Table S12. Measurement uncertainty using the Nordtest approach. Optical analyzer.
[7 empty] | Cresult-Cref (%) |  |  |  | Stirring test
HACH | [6 empty] | drift1 | drift 2 | temp_komp | act | sr | n
Possible bias: | [6 empty] | 0.2817% | 0.6672% | 0.1758% | 1.129% | 0.1099% | 30
Estimated by the data of drift estimation, temperature compensation tests and measurements in water with higher salinity. | [6 empty] | 1.283% | 0.5565% | 0.1211% | 0.9626% | 0.09733% | 30
Drift is measured two times during one month, alltogether 33 individual results. | [6 empty] | 1.356% | 0.3629% | -0.2511% | 1.49% | 0.07872% | 30
Temperature compensation tests include measurements in one day, alltogether 5 individual bias determinations. | [6 empty] | 1.432% | 0.1237% | -0.2398% | 1.842% | 0.08609% | 210
Measurements in KCl solution were made in one day and include 4 individual results. | [6 empty] | 0.6523% | -1.209% | -0.4075%
[7 empty] | 1.428% | -1.105%
RMSbias_rel | 1.311% | [5 empty] | 1.118% | -2.138%
u(Cref)rel | 0.4% | [5 empty] | 1.017% | -2.1%
[5 empty] | u(bias)= | 1.371% | 1.64% | -2.127%
Intermediate precision: | [6 empty] | 0.4918% | -2.162%
Estimated by the drift measurement data. | [6 empty] | 1.03% | -2.004%
Intermediate precision has been estimated as the pooled standard deviation of the relative differences of results from the reference value. | [6 empty] | 1.642% | -1.711%
Measurements were done with two different sensor calibration values whereby on calibration was used during one month. | [6 empty] | 1.081% | -2.328%
u(Rw)rel= | 0.2641% | [5 empty] | 0.8245% | -2.22%
[7 empty] | 0.9317% | -1.834%
[7 empty] | 0.3066% | -1.058%
uc= | 1.396% | [5 empty] | 0.8248%
So for the sample |  | 10 | mg/L | the standard uncertainty is | 0.13962 | mg/L
Note on B11: From the data in Ref 3

## Table S13
Table S13. Measurement uncertainty using the Nordtest approach. Amperometric analyzer.
[7 empty] | Cresult-Cref (%) |  |  |  | Stirring test
WTW | [6 empty] | drift1 | drift 2 | temp_komp | act | sr | n
Possible bias: | [6 empty] | 0.005548% | -0.007691% | -0.9002% | -0.2512% | 1.056% | 30
Estimated by the data of drift estimation, temperature compensation tests and measurements in water with higher salinity. | [6 empty] | -0.5626% | 0.1072% | -0.5559% | -0.3715% | 1.111% | 30
Drift is measured two times during one month, alltogether 30 individual results. | [6 empty] | -0.3817% | -0.556% | 0.4073% | 0.02398% | 1.099% | 30
Temperature compensation tests include measurements in one day, alltogether 5 individual bias determinations. | [6 empty] | -0.6871% | 0.09214% | 0.71% | -2.424%
Measurements in KCl solution were made in one day and include 4 individual results. | [6 empty] | -0.763% | 0.4395% | 0.9787%
[7 empty] | -0.9012% | 0.8358%
RMSbias_rel | 0.9718% |  |  |  | u(bias)= | 1.051% | 0.4043% | 0.03274%
u(Cref)rel | 0.4% | [5 empty] | -0.6862% | 0.4695%
[7 empty] | -1.598% | 0.4558%
Intermediate precision: | [6 empty] | -0.2762% | 0.6305%
Estimated by the drift measurement data. | [6 empty] | -0.5497% | 0.8669%
Intermediate precision has been estimated as the pooled standard deviation of the relative differences of results from the reference value. | [6 empty] | -0.4149% | 2.11%
Measurements were done with two different sensor calibration values whereby on calibration was used during one month. | [6 empty] | -1.682% | 2.617%
u(Rw)rel= | 1% | [5 empty] | -0.6119%
-1.101%
-1.286%
uc= | 1.451% | [5 empty] | -1.332%
So for the sample |  | 10 | mg/L | the standard uncertainty is | 0.14508 | mg/L
Note on B11: From the data in Ref 3

## References
References to the Electronic Supplementary Material:
1. HACH Company (2006). HQ Series Portable Meters, USER MANUAL, edition 5.
2. Weilheim 2002, WTW GmbH & Co. KG Operating Manual, CellOx 325 Dissolved oxygen sensor. Available at: http://www.globalw.com/downloads/WQ/cellox325.pdf
3. Helm, I., Jalukse, L., and Leito, I. (2012). A highly accurate method for determination of dissolved oxygen: Gravimetric Winkler method. Anal. Chim. Acta 741, 21–31. doi:10.1016/j.aca.2012.06.049.
4. Benson, B. B., and Krause, D. (1984). The concentration and isotopic fractionation of oxygen dissolved in freshwater and seawater in equilibrium with the atmosphere. Limnol. Oceanogr. 29, 620–632. doi:10.4319/lo.1984.29.3.0620.
5. McDougall, T.J. and P.M. Barker. Getting started with TEOS-10 and the Gibbs Seawater (GSW) Oceanographic Toolbox, 2011, 28pp., SCOR/IAPSO WG127, ISBN 978-0-646-55621-5.
6. Clesceri, L. S., American Public Health Association, American Water Works Association, and Water Pollution Control Federation eds. (1998). Standard methods: for the examination of water and wastewater. 20. ed. Washington: American Public Health Ass.
7. ISO 7888, Water quality - Determination of electrical conductivity, 1985.